## Sheet1
Output Characteristic Device | Output
Number of Output Channels1 - Synchronous, alternating - Method of channel isolation | 1, NA
Waveform2 (e.g., charge balanced biphasic symmetrical, biphasic asymmetrical) | Biphasic Symmetric
Pulse Shape (e.g., rectangular, sinusoidal) | Rectangular
Current/voltage regulated? Compliance voltage (if current source)? | Voltage Regulated
Maximum Output Voltage (specify units) (+/-______%) [voltage should be reported at 500 Ω and at impedances covering the minimum, typical, and maximum range of physiologic impedances for the location being stimulated] | 0.51
Maximum Output Current (specify units) (+/-______%) [current should be reported at 500 Ω and at impedances covering the minimum, typical, and maximum range of physiologic impedances for the location being stimulated] | 0.001
For multiphasic waveforms2 : - Symmetrical or Asymmetrical phases? - Phase Duration3 (include units) (state range, if applicable) (both phases, if asymmetrical) | Biphasic symmetrical
Pulse Duration2, 4 (specify units) | 0.00025
Frequency (Hz)5 | 30-240
Method of Balancing Charge6 | interleaved biphasic pulses
Are charge balancing cycles always completed? | Yes
Net Charge (µC per pulse) @ 500 Ω | Current / pulse width |  | Default
Leakage Current8 (nA) @ 500 Ω | Peak voltage / resistance | 647058.8235 | 240hz, 250 uS, 1 mA
Net DC Current9 (µA) at maximum pulse rate @ 500 Ω | Max volt peak/510- min volt/510 | 7.8431
Maximum Phase Charge (µC) @ 500 Ω | 2.6078e-07
Maximum Charge Density10 (µC/cm²/phase) @ 500 Ω
Maximum Phase Power (W/phase) @ 500 Ω | 0.000532
Maximum Phase Power Density (W/cm²/phase) @ 500 Ω | Use cell B17 and divide by elerode area
Pulse Delivery Mode (continuous/bursts (pulse trains)) | Burst
Burst Delivery11: a. Pulses per burst; b. Bursts per second; c. Burst duration (seconds); and d. Duty Cycle [Line a X Line b]
ON Time12 (seconds)
OFF Time12 (seconds)
Current Path Options13 (bipolar, unipolar, multipolar) | Bipolar
additional Features, if applicable | NA