## Sheet1
Output Characteristic | Output (software) | Output (measured) | Units
Number of Output Channels1 - Synchronous, alternating - Method of channel isolation | 1, NA | 1, NA
Waveform2 (e.g., charge balanced biphasic symmetrical, biphasic asymmetrical) | Biphasic Symmetric | Biphasic Symmetric
Pulse Shape (e.g., rectangular, sinusoidal) | Rectangular | Rectangular
Current/voltage regulated? Compliance voltage (if current source)? | Voltage Regulated | Voltage Regulated
Maximum Output Voltage (specify units) (+/-______%) [voltage should be reported at 500 Ω and at impedances covering the minimum, typical, and maximum range of physiologic impedances for the location being stimulated] | 0.51 | 0.51 | V | 524mV | v1 = 848, v2 = 806
Maximum Output Current (specify units) (+/-______%) [current should be reported at 500 Ω and at impedances covering the minimum, typical, and maximum range of physiologic impedances for the location being stimulated] | 0.001 | 0.001 | A
For multiphasic waveforms2 : - Symmetrical or Asymmetrical phases? - Phase Duration3 (include units) (state range, if applicable) (both phases, if asymmetrical) | Biphasic Symmetrical Phases | Biphasic Symmetrical Phases
Pulse Duration | 200 | 205.5 | us
Frequency (Stimulation Parameter 2) | 30-240 | 29.96-240 | Hz
Method of Balancing Charge | interleaved biphasic pulses | interleaved biphasic pulses
Are charge balancing cycles always completed? | Yes | Yes
Net Charge (µC per pulse) @ 510 Ω | 4 | 4 | uC |  | Default | Current / pulse width = 1ma / 250 us
Leakage Current (nA) @ 510 Ω | 647058.8235 | 647058.8235 | nA | 647058.8235 | 240hz, 250 uS, 1 mA | Peak voltage / resistance
Net DC Current (µA) at maximum pulse rate @ 510 Ω | 7.84 | 7.84 | uA | 7.8431 |  | Max volt peak/510- min volt/510
Maximum Phase Charge (µC) @ 510 Ω | 2.6078e-07 | 2.6078e-07 | uC
Maximum Charge Density10 (µC/cm²/phase) @ 510 Ω | 1.4765e-07 | 1.4765e-07
Maximum Phase Power (W/phase) @ 510 Ω | 0.000532 | 0.000532 | W/phase
Maximum Phase Power Density (W/cm²/phase) @ 510 Ω | 9.5303e-05 | 9.5303e-05 | W/cm^2/phase | Use cell B17 and divide by electrode area
Pulse Delivery Mode (continuous/bursts (pulse trains)) | Burst | Burst
Burst Delivery11: a. Pulses per burst; b. Bursts per second; c. Burst duration (seconds); and d. Duty Cycle [Line a X Line b] | Burst Delivery: a. Pulses per burst: 8.3 b. Bursts per second: 1 c. Burst duration (seconds): 1 d. Duty Cycle [Line a X Line b]: 8.3 | Burst Delivery: a. Pulses per burst: 8.3 b. Bursts per second: 1 c. Burst duration (seconds): 1 d. Duty Cycle [Line a X Line b]: 8.3 | seconds
ON Time(seconds) | 1 | 1 | seconds
OFF Time (seconds) | 1 | 1 | seconds
Current Path Options13 (bipolar, unipolar, multipolar) | Bipolar | Bipolar
additional Features, if applicable | NA | NA